## Sheet1
CALIBRATION CERTIFICATE FOR UNIVERSAL TESTING MACHINE
Calibration Certificate No. : |  |  |  |  | Date : 10 / 03 / 2005 |  |  |  | Page No. : 1 of 2
Name of Customer :
TECHNO FORGE LTD.
1022, G. I. D. C. ESTATE,
ANKLESHWAR,
Dist :- BHARUCH.
GUJARAT
Equipment Under Calibration: |  |  |  |  | Calibration Conditions:
Type |  | : UNIVERSAL TESTING M/c. |  |  | Referance Std. |  |  | : IS 1828 - 1991
Max. Capacity |  | : 40,000 Kgf |  |  | Cal.Procedure Ref. |  |  | : CP - 02
Make |  | : FIE |  |  | Date of Request |  |  | : 01 / 03 / 2005
Model |  | : UTN - 40 |  |  | Date of Calibration |  |  | : 10 / 03 / 2005
Sr. No. |  | : 9 / 93 - 1642 |  |  | Recommended Due Date |  |  | : 09 / 03 / 2006
I. D. No. |  | : ------ |  |  | Calibrated By |  |  | : NILESH PANCHAL
Location |  | : LABORATORY |  |  | Calibrated at Temp. |  |  | : 0 C
REFERANCE STANDARDS USED FOR LOAD VERIFICATION
|  |  | I |  |  | II
Make / Capacity |  |  | AWON / 100 KN |  |  | AWON / 400 KN
I. D. No. |  |  | KE / PR / 003 |  |  | KE / PR / 004
Sr. No. |  |  | 97506 |  |  | 97519
Gauge Sr. No. |  |  | G 3700 |  |  | J 2248
Calibration Cer. No. |  |  | FRI/03/04/1483 |  |  | FRI/03/04/1482
Calibration Due |  |  | 03 / 05 / 2006 |  |  | 03 / 05 / 2006
Calibration Temperature |  |  | 26 0 C |  |  | 26 0 C
Calibration Agency |  |  | FRI, ICHHALKARANJI |  |  | FRI, ICHHALKARANJI
NPL Traceability No. |  |  | 1028/04/347/2001 |  |  | 1028/04/347/2001
For Krutam Enterprises
Witness By: | [7 empty] | Calibrated By:
Mr. A.P.SADASHIVAM | [7 empty] | NILESH PANCHAL
L.R.A. - VODADARA | [7 empty] | (Sr. Cal. Engineer)
Measurements are Traceable to NPL
CALIBRATION CERTIFICATE FOR UNIVERSAL TESTING MACHINE
Calibration Certificate No. : |  |  |  |  | Date : 10 / 03 / 2005 |  |  |  | Page No. : 2 of 2
CALIBRATION RESULT
Proving | Load in | Std. P. R. | Temp. Crt. | Observed Reading (Div.) |  |  | Avg. | Diff. | Error | Rep.
Ring | Kgf | Reading | Standard | I | II | III | (Div.) | (Div.) | % | %
Used |  | (Div.) | (Div.)
Range I |  | 0 - 4000 Kgf | [5 empty] | L.C. :- 10 Kgf
100 KN | 1000 | 93.1 | 93.0246 |  |  |  | 0 | 0.2 | 0.215 | #DIV/0!
| 2000 | 187 | 186.8485 |  |  |  | 0 | 0.5 | 0.2676 | #DIV/0!
| 3000 | 281.4 | 281.1721 |  |  |  | 0 | -1.2 | -0.42678 | #DIV/0!
| 3500 | 328.7 | 328.4338 |  |  |  | 0 | -1.9 | -0.5785 | #DIV/0!
| 4000 | 376.2 | 375.8953 |  |  |  | 0 | -0.9 | -0.23943 | #DIV/0!
Range II |  | 0 - 10000 Kgf | [5 empty] | L.C. :- 25 Kgf
100 KN | 2000 | 187 | 186.8485 |  |  |  | 0 | -0.1 | -0.053519 | #DIV/0!
| 4000 | 376.2 | 375.8953 |  |  |  | 0 | 0.1 | 0.026603 | #DIV/0!
| 6000 | 567.1 | 566.6406 |  |  |  | 0 | -0.6 | -0.10589 | #DIV/0!
| 8000 | 759.7 | 759.0846 |  |  |  | 0 | -2.8 | -0.36887 | #DIV/0!
| 10000 | 953.9 | 953.1273 |  |  |  | 0 | -6.8 | -0.71344 | #DIV/0!
Range III |  | 0 - 20000 Kgf | [5 empty] | L.C. :- 50 Kgf
400 KN | 4000 | 62.3 | 62.2495 |  |  |  | 0 | 0.3 | 0.48193 | #DIV/0!
| 8000 | 125.8 | 125.6981 |  |  |  | 0 | 1.1 | 0.87511 | #DIV/0!
| 12000 | 188.9 | 188.747 |  |  |  | 0 | 0.7 | 0.37087 | #DIV/0!
| 16000 | 251.8 | 251.596 |  |  |  | 0 | 0.7 | 0.27822 | #DIV/0!
| 20000 | 314.7 | 314.4451 |  |  |  | 0 | 2.6 | 0.82685 | #DIV/0!
Range IV |  | 0 - 40000 Kgf | [5 empty] | L.C. :- 100 Kgf
400 KN | 8000 | 125.8 | 125.6981 |  |  |  | 0 | 1.2 | 0.95467 | #DIV/0!
| 16000 | 251.8 | 251.596 |  |  |  | 0 | 1.4 | 0.55645 | #DIV/0!
| 24000 | 377.6 | 377.2941 |  |  |  | 0 | 1.5 | 0.39757 | #DIV/0!
| 32000 | 504.7 | 504.2912 |  |  |  | 0 | 1.7 | 0.33711 | #DIV/0!
| 40000 | 634.2 | 633.6863 |  |  |  | 0 | -0.4 | -0.063123 | #DIV/0!
Note: - Error allowed is ± 1% temperature corrected standard
Allowed Repeatability is ± 1% of temperature corrected standard
Grade of the machine – Class 1 Machine
For Krutam Enterprises
Witness By: | [7 empty] | Calibrated By:
Mr. A.P.SADASHIVAM | [7 empty] | NILESH PANCHAL
L.R.A. - VODADARA | [7 empty] | (Sr. Cal. Engineer)
1. The above calibration is done at the site of customer.
2. The calibration results reported in this certificate are valid at the time of and under the
stated conditions of measurements and is only for the calibrated item as identified in the
certificate.
3. Calibration certificate shall not be reproduced except in full, without written approval
of KRUTAM ENTERPRISES

## Conversion Chart
HARDNESS CONVERSION CHART
HARDENED STEEL AND HARD ALLOYS | [10 empty] | Soft steel, grey and malleable cast iron and most non-ferrous metals
ROCKWELL STANDARD |  |  | ROCKWELL SUPERFICIAL |  |  | VICKERS | KNOOP | BRINELL | TENSILE STRENGTH N / mm2 | SHORE SCLEROSCOPE HARDNESS NUMBER | ROCKWELL STANDARD |  | ROCKWELL SUPERFICIAL |  |  | ROCKWELL STANDARD |  |  |  | KNOOP | BRINELL | BIRNELL VICKERS | TENSILE STRENGTH N / mm2
C | A | D | 15 - N | 30 - N | 45 - N | HV | HK | BHN |  |  | B | F | 15 - T | 30 - T | 45 - T | E | H | K | A | HK | HB | BHN, HV
150 Kg | 60 Kg | 100 Kg | 15 Kg | 30 Kg | 45 Kg | 10 Kg | 500 gm & Above | 3000 Kg 10 mm Ball |  |  | 100 Kg | 60 Kg | 15 Kg | 30 Kg | 45 Kg | 100 Kg | 60 Kg | 150 Kg | 60 Kg Diamond | 500 gm & Above | 500 Kg 10 mm Ball | 3000 Kg 10 mm Ball 10 Kg
DIAMOND INDENTOR |  |  | DIAMOND INDENTOR | [7 empty] | 1/16" Ball |  | 1/16" Ball |  |  | 1 / 8" Ball
80 | 92 | 86.5 | 96.5 | 92 | 87 | 1865 |  |  |  |  | 100 |  | 93.1 | 83.1 | 72.9 |  |  | 100 | 61.5 | 251 | 201 | 240 | 799
79 | 91.5 | 85.5 | 96.3 | 91.5 | 86.5 | 1787 |  |  |  |  | 99 |  | 92.8 | 82.5 | 71.9 |  |  |  | 60.9 | 246 | 195 | 234 | 785
78 | 91 | 84.5 | 96 | 91 | 85.5 | 1710 |  |  |  |  | 98 |  | 92.5 | 81.8 | 70.9 |  |  |  | 60.2 | 241 | 189 | 228 | 751
77 | 90.5 | 84 | 95.8 | 90.5 | 84.5 | 1633 |  |  |  |  | 97 |  | 92.1 | 81.1 | 69.9 |  |  |  | 59.5 | 236 | 184 | 222 | 717
76 | 90 | 83 | 95.5 | 90 | 83.5 | 1556 |  |  |  |  | 96 |  | 91.8 | 80.4 | 68.9 |  |  |  | 58.9 | 231 | 179 | 216 | 703
75 | 89.5 | 82.5 | 95.3 | 89 | 82.5 | 1478 | NA |  |  | NA | 95 |  | 91.5 | 79.8 | 67.9 |  |  |  | 58.3 | 226 | 175 | 210 | 689
74 | 89 | 81.5 | 95 | 88.5 | 81.5 | 1400 |  |  |  |  | 94 |  | 91.2 | 79.1 | 66.9 |  |  |  | 57.6 | 221 | 171 | 205 | 675
73 | 88.5 | 81 | 94.8 | 88 | 80.5 | 1323 |  | NA |  |  | 93 |  | 90.8 | 78.4 | 65.9 |  |  |  | 57 | 216 | 167 | 200 | 648
72 | 88 | 80 | 94.5 | 87 | 79.5 | 1245 |  |  |  |  | 92 |  | 90.5 | 77.8 | 64.8 |  |  |  | 56.4 | 211 | 163 | 195 | 634
71 | 87 | 79.5 | 94.3 | 86.5 | 78.5 | 1160 |  |  | NA |  | 91 |  | 90.2 | 77.1 | 63.8 |  |  | 99.5 | 55.8 | 206 | 160 | 190 | 620
70 | 86.5 | 78.5 | 94 | 86 | 77.5 | 1076 | 972 |  |  |  | 90 | NA | 89.9 | 76.4 | 62.8 |  |  | 98.5 | 55.2 | 201 | 157 | 185 | 613
69 | 86 | 78 | 93.5 | 85 | 76.5 | 1004 | 946 |  |  |  | 89 |  | 89.5 | 75.8 | 61.8 |  |  | 98 | 54.6 | 196 | 154 | 180 | 606
68 | 85.6 | 76.9 | 93.2 | 84.4 | 75.4 | 940 | 920 |  |  | 97 | 88 |  | 89.2 | 75.1 | 60.8 |  |  | 97 | 54 | 192 | 151 | 176 | 593
67 | 85 | 76.1 | 92.9 | 83.6 | 74.2 | 900 | 895 |  |  | 95 | 87 |  | 88.9 | 74.4 | 59.8 | NA |  | 96.5 | 53.4 | 188 | 148 | 172 | 579
66 | 84.5 | 75.4 | 92.5 | 82.8 | 73.3 | 865 | 870 |  |  | 92 | 86 |  | 88.6 | 73.8 | 58.8 |  |  | 95.5 | 52.8 | 184 | 145 | 169 | 572
65 | 83.9 | 74.5 | 92.2 | 81.9 | 72 | 832 | 846 | 739 |  | 91 | 85 |  | 88.2 | 73.1 | 57.8 |  |  | 94.5 | 52.3 | 180 | 142 | 165 | 565
64 | 83.4 | 73.8 | 91.8 | 81.1 | 71 | 800 | 822 | 722 |  | 88 | 84 |  | 87.9 | 72.4 | 56.8 |  |  | 94 | 51.7 | 176 | 140 | 162 | 558
63 | 82.8 | 73 | 91.4 | 80.1 | 69.9 | 772 | 799 | 706 |  | 87 | 83 |  | 87.6 | 71.8 | 55.8 |  |  | 93 | 51.1 | 173 | 137 | 159 | 551
62 | 82.3 | 72.2 | 91.1 | 79.3 | 68.8 | 746 | 776 | 688 |  | 85 | 82 |  | 87.3 | 71.1 | 54.8 |  |  | 92 | 50.6 | 170 | 135 | 156 | 531
61 | 81.8 | 71.5 | 90.7 | 78.4 | 67.7 | 720 | 754 | 670 |  | 83 | 81 |  | 86.9 | 70.4 | 53.8 |  |  | 91 | 50 | 167 | 133 | 153 | 503
60 | 81.2 | 70.7 | 90.2 | 77.5 | 66.6 | 697 | 732 | 654 |  | 81 | 80 |  | 86.6 | 69.7 | 52.8 |  |  | 90.5 | 49.5 | 164 | 130 | 150 | 496
59 | 80.7 | 69.9 | 89.8 | 76.6 | 65.5 | 674 | 710 | 634 | 2418 | 80 | 79 |  | 86.3 | 69.1 | 51.8 |  |  | 89.5 | 48.9 | 161 | 128 | 147 | 482
58 | 80.1 | 69.2 | 89.3 | 75.7 | 64.3 | 653 | 690 | 615 | 2329 | 78 | 78 |  | 86 | 68.4 | 50.8 |  |  | 88.5 | 48.4 | 158 | 126 | 144 | 475
57 | 79.6 | 68.5 | 88.9 | 74.8 | 63.2 | 633 | 670 | 595 | 2239 | 76 | 77 |  | 85.6 | 67.7 | 49.8 |  | NA | 88 | 47.9 | 155 | 124 | 141 | 467
56 | 79 | 67.7 | 88.3 | 73.9 | 62 | 613 | 650 | 577 | 2156 | 75 | 76 |  | 85.3 | 67.1 | 48.8 |  |  | 87 | 47.3 | 152 | 122 | 139 | 462
55 | 78.5 | 66.9 | 87.9 | 73 | 60.9 | 595 | 630 | 560 | 2074 | 74 | 75 | 99.6 | 85 | 66.4 | 47.8 |  |  | 86 | 46.8 | 150 | 120 | 137 | 455
54 | 78 | 66.1 | 87.4 | 72 | 59.6 | 577 | 612 | 543 | 2012 | 72 | 74 | 99.1 | 84.7 | 65.7 | 46.8 |  |  | 85 | 46.3 | 147 | 118 | 135 | 448
53 | 77.4 | 65.4 | 86.9 | 71.2 | 58.6 | 560 | 594 | 525 | 1950 | 71 | 73 | 98.5 | 84.3 | 65.1 | 45.8 |  |  | 84.5 | 45.8 | 145 | 116 | 132 | 441
52 | 76.8 | 64.6 | 86.4 | 72 | 57.4 | 544 | 576 | 512 | 1881 | 69 | 72 | 98 | 84 | 64.4 | 44.8 |  |  | 83.5 | 45.3 | 143 | 114 | 130 | 434
51 | 76.3 | 63.8 | 85.9 | 69.4 | 56.1 | 528 | 558 | 496 | 1819 | 68 | 71 | 97.7 | 83.7 | 63.7 | 43.8 | 100 |  | 82.5 | 44.8 | 141 | 112 | 127 | 427
50 | 75.9 | 63.1 | 85.5 | 68.5 | 55 | 513 | 542 | 481 | 1757 | 67 | 70 | 96.8 | 83.4 | 63.1 | 42.8 | 99.5 |  | 81.5 | 44.3 | 139 | 110 | 125 | 420
49 | 75.2 | 62.1 | 85 | 67.6 | 53.8 | 498 | 526 | 469 | 1695 | 66 | 69 | 96.2 | 83 | 62.4 | 41.8 | 99 |  | 81 | 43.8 | 137 | 109 | 132 | 413
48 | 74.7 | 61.4 | 84.5 | 66.7 | 52.5 | 484 | 510 | 455 | 1640 | 64 | 68 | 95.6 | 82.7 | 61.7 | 40.8 | 98 |  | 80 | 43.3 | 135 | 107 | 121 | 407
47 | 74.1 | 60.8 | 83.9 | 65.8 | 51.4 | 471 | 495 | 443 | 1578 | 63 | 67 | 95.1 | 82.4 | 61 | 39.8 | 97.5 |  | 79 | 42.8 | 133 | 106 | 119 | 400
46 | 73.6 | 60 | 83.5 | 64.8 | 50.3 | 458 | 480 | 432 | 1523 | 62 | 66 | 94.5 | 82.1 | 60.4 | 38.7 | 97 |  | 78.5 | 42.3 | 131 | 104 | 117 | 393
45 | 73.1 | 59.2 | 83 | 64 | 49 | 446 | 466 | 421 | 1481 | 60 | 65 | 93.9 | 81.8 | 59.7 | 37.7 | 96 |  | 77.5 | 41.8 | 129 | 102 | 116 | 386
44 | 72.5 | 58.5 | 82.5 | 63.1 | 47.8 | 434 | 452 | 409 | 1433 | 58 | 64 | 93.4 | 81.4 | 59 | 36.7 | 95.5 |  | 76 | 41.4 | 127 | 101 | 114
43 | 72 | 57.7 | 82 | 62.2 | 46.7 | 423 | 438 | 400 | 1385 | 57 | 63 | 92.8 | 81.1 | 58.4 | 35.7 | 95 |  | 74.5 | 40.9 | 125 | 99 | 112
42 | 71.5 | 56.9 | 81.5 | 61.3 | 45.5 | 412 | 426 | 390 | 1337 | 56 | 62 | 90.2 | 80.8 | 57.7 | 34.7 | 94.5 |  | 74 | 40.4 | 124 | 98 | 110
41 | 70.9 | 56.2 | 80.9 | 60.4 | 44.3 | 402 | 414 | 381 | 1295 | 55 | 61 | 91.7 | 80.5 | 57 | 33.7 | 93.5 |  | 74 | 40 | 122 | 96 | 108
40 | 70.4 | 55.4 | 80.4 | 59.5 | 43.1 | 392 | 402 | 371 | 1254 | 54 | 60 | 91.1 | 80.1 | 56.4 | 32.7 | 93 |  | 73 | 39.5 | 120 | 95 | 107
39 | 69.9 | 54.6 | 79.9 | 58.8 | 41.9 | 382 | 391 | 362 | 1220 | 52 | 59 | 90.5 | 79.8 | 55.7 | 31.7 | 92.5 |  | 72 | 39 | 118 | 94 | 106
38 | 69.4 | 53.8 | 79.4 | 57.7 | 40.8 | 372 | 380 | 353 | 1178 | 51 | 58 | 90 | 75.5 | 55 | 30.7 | 92 |  | 71 | 38.6 | 117 | 92 | 104
37 | 68.9 | 53.1 | 78.8 | 56.8 | 39.6 | 363 | 370 | 344 | 1144 | 50 | 57 | 89.4 | 79.2 | 54.4 | 29.7 | 91 |  | 70.5 | 38.1 | 115 | 91 | 103
36 | 68.4 | 52.3 | 78.3 | 55.9 | 38.4 | 354 | 360 | 336 | 1109 | 49 | 56 | 88.8 | 78.8 | 53.7 | 28.7 | 90.5 |  | 69.5 | 37.7 | 114 | 90 | 101
35 | 67.9 | 51.5 | 77.7 | 55 | 37.2 | 345 | 351 | 327 | 1075 | 48 | 55 | 88.2 | 78.5 | 53 | 27.7 | 90 |  | 68.5 | 37.2 | 112 | 89 | 100
34 | 67.4 | 50.8 | 77.2 | 54.2 | 36.1 | 336 | 342 | 319 | 1047 | 47 | 54 | 87.7 | 78.2 | 52.4 | 26.7 | 89.5 |  | 68 | 36.8 | 111 | 87
33 | 66.8 | 50 | 76.6 | 53.3 | 34.9 | 327 | 334 | 311 | 1027 | 46 | 53 | 87.1 | 77.9 | 51.7 | 25.7 | 89 |  | 67 | 36.3 | 110 | 86
32 | 66.3 | 49.2 | 76.1 | 52.1 | 33.7 | 318 | 326 | 301 | 1006 | 45 | 52 | 86.5 | 77.5 | 51 | 24.7 | 88 |  | 66 | 35.9 | 109 | 85
31 | 65.8 | 48.4 | 75.6 | 51.3 | 32.5 | 310 | 318 | 294 | 971 | 44 | 51 | 86 | 77.2 | 50.3 | 23.7 | 87.5 |  | 65 | 35.5 | 108 | 84
30 | 65.3 | 47.7 | 75 | 50.4 | 31.3 | 302 | 311 | 286 | 951 | 43 | 50 | 85.4 | 76.9 | 49.7 | 22.7 | 87 |  | 64.5 | 35 | 107 | 83
29 | 64.6 | 47 | 74.5 | 49.5 | 30.1 | 294 | 304 | 279 | 930 | 42 | 49 | 84.8 | 76.6 | 49 | 21.7 | 86.5 |  | 63.5 | 34.6 | 106 | 82
28 | 64.3 | 46.1 | 73.9 | 48.6 | 28.9 | 286 | 297 | 271 | 903 | 41 | 48 | 84.3 | 76.2 | 48.3 | 20.7 | 85.5 |  | 62.5 | 34.1 | 105 | 81
27 | 63.8 | 45.2 | 73.3 | 47.7 | 27.8 | 279 | 290 | 264 | 882 | 40 | 47 | 83.7 | 75.9 | 47.7 | 19.7 | 85 |  | 61.5 | 33.7 | 104 | 80
| 63.3 | 44.6 | 72.8 | 46.8 | 26.7 | 272 | 284 | 258 | 861 | 39 | 46 | 83.1 | 75.6 | 47 | 18.7 | 84.5 |  | 61 | 33.3 | 103 | 80
25 | 62.8 | 43.8 | 72.2 | 45.9 | 25.5 | 266 | 278 | 253 | 847 | 38 | 45 | 82.6 | 75.3 | 46.3 | 17.7 | 84 |  | 60 | 32.9 | 102 | 79
24 | 62.4 | 43.1 | 71.6 | 45 | 24.3 | 260 | 272 | 247 | 820 | 37 | 44 | 82 | 74.9 | 45.7 | 16.7 | 83.5 |  | 59 | 32.4 | 101 | 78
23 | 60 | 42.1 | 71 | 44 | 23.1 | 254 | 266 | 243 | 806 | 36 | 43 | 81.4 | 74.6 | 45 | 15.7 | 82.5 | NA | 58 | 32 | 100 | 77
22 | 61.1 | 41.6 | 70.5 | 43.2 | 22 | 248 | 261 | 237 | 792 | 35 | 42 | 80.8 | 74.3 | 44.3 | 14.7 | 82 |  | 57.5 | 31.6 | 99 | 76
21 | 61 | 40.9 | 69.9 | 42.3 | 20.7 | 243 | 256 | 231 | 772 | 34 | 41 | 80.3 | 74 | 43.7 | 13.6 | 81.5 |  | 56.5 | 31.2 | 98 | 75
20 | 60.5 | 40.1 | 69.4 | 41.5 | 19.6 | 238 | 251 | 226 | 758 | 33 | 40 | 79.7 | 73.6 | 43 | 12.6 | 81 |  | 55.5 | 30.7 | 97 | 75
[11 empty] | 39 | 79.1 | 73.3 | 42.3 | 11.6 | 80 |  | 54.5 | 30.3 | 96 | 74
[11 empty] | 38 | 78.6 | 73 | 41.6 | 10.6 | 79.5 |  | 54 | 29.9 | 95 | 73
[11 empty] | 37 | 78 | 72.7 | 41 | 9.6 | 79 |  | 53 | 29.5 | 94 | 72
[11 empty] | 36 | 77.4 | 72.3 | 40.3 | 8.6 | 78.5 | 100 | 52 | 29.1 | 93 | 72
CYLINDRICAL CORRECTION | [10 empty] | 35 | 76.9 | 72 | 39.6 | 7.6 | 78 | 99.5 | 51.5 | 28.7 | 92 | 71
to be added to observed Rockwell Number for Scales Indicated | [10 empty] | 34 | 76.3 | 71.7 | 39 | 6.6 | 77 | 99 | 50.5 | 28.2 | 91 | 70
SCALES C,D,A | [10 empty] | 33 | 75.7 | 71.4 | 38.3 | 5.6 | 76.5 | 98.8 | 49.5 | 27.8 | 90 | 69
DIAMOND INDENTOR |  |  |  |  | DIAMETER OF SPECIMEN - mm | [5 empty] | 32 | 75.2 | 71 | 37.6 | 4.6 | 76 | 98.5 | 48.5 | 27.4 | 89 | 69 |  | NA
Observed Reading | 3 | 6 | 10 | 13 | 16 | 19 | 22 | 25 |  |  | 31 | 74.6 | 70.7 | 37 | 3.6 | 75.5 | 98 | 48 | 27 | 88 | 68
90 | N/A | 0.5 | 0 | 0 | 0 | 0 | 0 | 0 |  |  | 30 | 74 | 70.4 | 36.3 | 2.6 | 75 | 97.8 | 47 | 26.6 | 87 | 67
85 |  | 0.5 | 0.5 | 0.5 | 0 | 0 | 0 | 0 |  |  | 29 | 73.5 | 70 | 35.6 | 1 | 74 | 97.5 | 46 | 26 | 87 | 66
80 |  | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0 | 0 |  |  | 28 | 73 | 69.5 | 34.5 |  | 73.5 | 97 | 45 | 25.5 | 86 | 66 | NA
75 |  | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 | 0 |  |  | 26 | 72 | 69 | 33 |  | 72.5 | 96.3 | 43.5 | 24.5 | 84 | 65
70 |  | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 25 | 71 | 68.8 | 32.5 |  | 72 | 96 | 42.5 | 24.3 | 83 | 64
65 |  | 1.5 | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 24 | 70.5 | 68.5 | 32 |  | 71 | 95.5 | 41.5 | 24 | 82 | 64
60 |  | 1.5 | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 23 | 70 | 68 | 31 |  | 70.5 | 95.3 | 41 | 23.5 | 82 | 63
55 |  | 2 | 1.5 | 1 | 1 | 0.5 | 0.5 | 0.5 |  |  | 22 | 69.5 | 67.8 | 30.5 |  | 70 | 95 | 40 | 23 | 81 | 62
50 |  | 2.5 | 2 | 1.5 | 1 | 1 | 0.5 | 0.5 |  |  | 21 | 69 | 67.5 | 29.5 |  | 69.5 | 94.5 | 39 | 22.5 | 81 | 62
45 |  | 3 | 2 | 1.5 | 1 | 1 | 1 | 0.5 |  |  | 20 | 68.5 | 67.3 | 29 |  | 68.5 | 94.3 | 38 | 22 | 80 | 61
40 |  | 3.5 | 2.5 | 2 | 1.5 | 1 | 1 | 1 |  |  | 19 | 68 | 67 | 28.5 |  | 68 | 94 | 37.5 | 21.5 | 79 | 61
35 |  | 4 | 3 | 2 | 1.5 | 1.5 | 1 | 1 |  |  | 18 | 67 | 66.5 | 27.5 |  | 67.5 | 93.5 | 36.5 | 21.3 | 78 | 60
30 |  | 5 | 3.5 | 2.5 | 2 | 1.5 | 1.5 | 1 |  |  | 17 | 66.5 | 66.3 | 27 |  | 67 | 93 | 35.5 | 21 | 78 | 60
25 |  | 5.5 | 4 | 3 | 2.5 | 2 | 1.5 | 1 |  |  | 16 | 66 | 66 | 26 |  | 66 | 92.8 | 35 | 20.5 | 77 | 59
20 |  | 6 | 4.5 | 3.5 | 2.5 | 2 | 1.5 | 1.5 |  |  | 15 | 65.5 | 65.5 | 25.5 |  | 65.5 | 92.5 | 34 | 20 | 76 | 59
SCALES B,F | [10 empty] | 14 | 65 | 65.3 | 25 | NA | 65 | 92 | 33 |  | 75 | 59
1/16" Ball Indentor |  |  |  |  | DIAMETER OF SPECIMEN - mm | [5 empty] | 13 | 64.5 | 65 | 24 |  | 64.5 | 91.8 | 32 |  | 75 | 58
Observed Reading | 3 | 6 | 10 | 13 | 16 | 19 | 22 | 25 |  |  | 12 | 64 | 64.5 | 23.5 |  | 64 | 91.5 | 31.5 |  | 74 | 58
[11 empty] | 11 | 63.5 | 64.3 | 23 |  | 63.5 | 91 | 30.5 |  | 73 | 57
100 | N/A | 3.5 | 2.5 | 1.5 | 1.5 | 1 | 1 | 0.5 |  |  | 10 | 63 | 64 | 22 |  | 62.5 | 90.5 | 29.5 |  | 72 | 57
90 |  | 4 | 3 | 2 | 1.5 | 1.5 | 1.5 | 1 |  |  | 9 | 62 | 63.8 | 21.5 |  | 62 | 90.3 | 29 |  | 71 | 57
80 |  | 5 | 3.5 | 2.5 | 2 | 1.5 | 1.5 | 1.5 |  |  | 8 | 61.5 | 63.5 | 20.5 |  | 61.5 | 90 | 28 |  | 71 | 56
70 |  | 6 | 4 | 3 | 2.5 | 2 | 2 | 1.5 |  |  | 7 | 61 | 63 | 20 |  | 61 | 89.5 | 27 | NA | 70 | 56
60 |  | 7 | 5 | 3.5 | 3 | 2.5 | 2 | 2 |  |  | 6 | 60.5 | 62.8 | 19.5 |  | 60.5 | 89.3 | 26 |  | 69 | 55
50 |  | 8 | 5.5 | 4 | 3.5 | 3 | 2.5 | 2 |  |  | 5 | 60 | 62.5 | 18.5 |  | 60 | 89 | 25.5 |  | 69 | 55
40 |  | 9 | 6 | 4.5 | 4 | 3 | 2.5 | 2.5 |  |  | 4 | 59.5 | 62 | 18 |  | 59 | 88.5 | 24.5 |  | 69 | 55
30 |  | 10 | 6.5 | 5 | 4.5 | 3.5 | 3 | 2.5 |  |  | 3 | 59 | 61.8 | 17 |  | 58.5 | 88 | 23.5 |  | 68 | 54
20 |  | 11 | 7.5 | 5.5 | 4.5 | 4 | 3.5 | 3 |  |  | 2 | 58 | 61.5 | 16.5 |  | 58 | 87.8 | 23 |  | 68 | 54
10 |  | 12 | 8 | 6 | 5 | 4 | 3.5 | 3 |  |  | 1 | 57.5 | 61 | 16 |  | 57.5 | 87.5 | 22 |  | 68 | 54
0 |  | 12.5 | 8.5 | 6.5 | 5.5 | 4.5 | 3.5 | 3 |  |  | 0 | ----- | ----- | ----- |  | ----- | ----- | 21 |  | ----- | -----
SCALES 15-N,30-N, 45-N
DIAMOND INDENTOR |  |  |  |  | DIAMETER OF SPECIMEN - mm
Observed Reading | 3 | 6 | 10 | 13 | 16 | 19 | 22 | 25 |  |  | HARDENSS Vs MINIMUM THICKNESS
[11 empty] | Any greater thickness and hardness can be Safely tested on Indicated Scale |  |  | Rockwell Superficial Hardness Scales |  |  | Rockwell Standard Hardenss Scales
90 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
85 | 0.5 | 0.5 | 0.5 | 0.5 | 0 | 0 | 0 | 0 | [5 empty] | 15 N | 30 N | 45 N | A | D | C
80 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0 | 0 | 0 | [5 empty] | 15 Kgf | 30 Kgf | 45 Kgf | 60 Kgf | 10 Kgf | 150 Kgf
75 | 1.5 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0 | 0 |  |  | Thickness (mm) |  |  | Diamond Indentor |  |  | Diamond Indentor
70 | 2 | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 0.15 |  |  | 92 | - | - | - | - | -
65 | 2.5 | 1.5 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 0.2 |  |  | 90 | - | - | - | - | -
60 | 3 | 1.5 | 1 | 1 | 1 | 0.5 | 0.5 | 0.5 |  |  | 0.25 |  |  | 88 | - | - | - | - | -
55 | 3.5 | 2 | 1.5 | 1 | 1 | 0.5 | 0.5 | 0.5 |  |  | 0.3 |  |  | 80 | 62 | - | - | - | -
50 | 3.5 | 2 | 1.5 | 1 | 1 | 1 | 1 | 0.5 |  |  | 0.35 |  |  | 76 | 78.5 | - | - | - | -
45 | 4 | 2 | 1.5 | 1 | 1 | 1 | 1 | 1 |  |  | 0.4 |  |  | 68 | 74 |  | 83 | - | -
40 | 4.5 | 2.5 | 1.5 | 1.5 | 1 | 1 | 1 | 1 |  |  | 0.45 |  |  | ü | 66 | 68 | 84 | - | -
35 | 5 | 2.5 | 2 | 1.5 | 1 | 1 | 1 | 1 |  |  | 0.5 |  |  | ü | 57 | 63 | 82 | 77 | -
30 | 5.5 | 3 | 2 | 1.5 | 1.5 | 1 | 1 | 1 |  |  | 0.55 |  |  | ü | 47 | 58 | 79 | 75 | 69
25 | 5.5 | 3 | 2 | 1.5 | 1.5 | 1.5 | 1.5 | 1 |  |  | 0.6 |  |  | ü | ü | 51 | 76 | 72 | 67
20 | 6 | 3 | 2 | 1.5 | 1.5 | 1.5 | 1.5 | 1.5 |  |  | 0.65 |  |  | ü | ü | 37 | 71 | 68 | 65
SCALES 15-T,30-T,45-T | [10 empty] | 7 |  |  | ü | ü | 20 | 67 | 63 | 62
1/16" Ball Indentor |  |  |  |  | DIAMETER OF SPECIMEN - mm | [5 empty] | 0.75 |  |  | ü | ü | ü | 60 | 58 | 57
Observed Reading | 3 | 6 | 10 | 13 | 16 | 19 | 22 | 25 |  |  | 0.8 |  |  | ü | ü | ü |  | 51 | 52
[11 empty] | 0.85 |  |  | ü | ü | ü | ü | 43 | 45
90 | 1.5 | 1 | 1 | 0.5 | 0.5 | 0.5 | 0.5 | 0.5 |  |  | 0.9 |  |  | ü | ü | ü | ü | ü | 37
80 | 3 | 2 | 1.5 | 1.5 | 1 | 1 | 1 | 0.5 |  |  | 0.95 |  |  | ü | ü |  | ü | ü | 28
70 | 5 | 3.5 | 2.5 | 2 | 1.5 | 1 | 1 | 1 |  |  | 1 |  |  | ü | ü | ü | ü | ü | 20
60 | 6.5 | 4.5 | 3 | 2.5 | 2 | 1.5 | 1 | 1.5 |  |  | Any greater thickness and hardness can be Safely tested on Indicated Scale |  |  | Rockwell Superficial Hardness Scales |  |  | Rockwell Standard Hardenss Scales
50 | 8.5 | 5.5 | 4 | 3 | 2.5 | 2 | 1.5 | 1.5
40 | 10 | 6.5 | 4.5 | 3.5 | 3 | 2.5 | 2 | 2 | [5 empty] | 15 T | 30 T | 45 T | F | B | G
30 | 11.5 | 7.5 | 5 | 3.5 | 3.5 | 2.5 | 2 | 2 | [5 empty] | 15 Kgf | 30 Kgf | 45 Kgf | 60 Kgf | 10 Kgf | 150 Kgf
20 | 13 | 9 | 6 | 4.5 | 4.5 | 0.3 | 2 | 2 |  |  | Thickness (mm) |  |  | 1/16" Ball Indentor |  |  | 1/16" Ball Indentor
[11 empty] | 0.25 |  |  | 91 | - | - | - | - | -
These Correction values are added to the dial gauge reading when hardness testing on the outer (Convex) surface and subtracted when testing on the inner (concave) surface. The values are only approximate and represent the averages to the nearest half Rockwell number. When testing cylindrical specimens, the accuracy of the test will be seriously affected by alignment of elevating screw, Vee anvil, Indentors, Surface finish and the straightness of the cylinder. | [10 empty] | 0.3 |  |  | 86 | - | - | - | - | -
[11 empty] | 0.35 |  |  | 81 | 80 | - | - | - | -
[11 empty] | 0.4 |  |  | 75 | 72 | 71 | - | - | -
[11 empty] | 0.45 |  |  | 68 | 64 | 62 | - | - | -
[11 empty] | 0.5 |  |  | ü | 55 | 53 | - | - | -
[11 empty] | 0.55 |  |  | ü | 45 | 43 | - | - | -
[11 empty] | 0.6 |  |  | ü | 34 | 31 | 98 | 94 | 94
[11 empty] | 0.65 |  |  | ü | ü | 18 | 91 | 87 | 87
[11 empty] | 0.7 |  |  | ü | ü | ü | 85 | 80 | 76
[11 empty] | 0.75 |  |  | ü | ü | ü | 77 | 71 | 68
[11 empty] | 0.8 |  |  | ü | ü | ü | 69 | 62 | 59
[11 empty] | 0.85 |  |  | ü | ü | ü | ü | 52 | 50
[11 empty] | 0.9 |  |  | ü | ü | ü | ü | 40 | 42
[11 empty] | 0.95 |  |  | ü | ü | ü | ü | 28 | 31
[11 empty] | 1 |  |  | ü | ü | ü | ü | ü | 22
ü No Minimum Hardness Applicable at this Thickness
Hardness conversions are empirical relationship and hence the above table is only approximate. From a practical stand point, this is of considerable value to relate different hardness scales. The above table assumes that the material tested is homogenous to a depth many times as great as the depth of the indentation. The indentation hardness values observed on the various scales depend on the work hardening behaviour of the material during the test. This depends on previous cold working and other process carried out on the material.
The Rockwell "B" Scale relationships are based mainly on annealed metals for the low values and cold worked metals for the higher values. Hence annealed austenite steels, Nickel and high Nickel alloys etd., do not confirm closely to this general tables. Similarly cold worked aluminium alloys and other soft alloys with low "B" scale value also do not correlate well with this table. For specific cases, we recommend more exact relations to be derived on a case to case basis. The tensile values again are very approximate. Important : Steel Ball is used for Brinell values SS than 450BHN. To measure more than 450BHN, Carbide Balls must be used.